| NO. | RANGE AND UNCERTAINTY | METHOD OF CALIBRATION | TYPES OF EQUIPEMENT TO BE CALIBRATED | TYPE OF STANDARDS USED BY BOM |
| --- | --- | --- | --- | --- |
| 1 | 100mV-1kV U> 3 ppm k =2 | LP-04-E-01 Laboratory procedure for calibration of digital multimeters | Multimeters/DC Voltmeters | Calibrator Wavetek 4808 |
| 2 | (100mV-100V)@( 10Hz-1MHz) U> 30 ppm k =2 | LP-04-E-01 Laboratory procedure for calibration of digital multimeters | Multimeters/AC Voltmeters | Calibrator Wavetek 4808 |
| 3 | (100V-1kV)@( 10Hz-100kHz) U> 100 ppm k =2 | LP-04-E-01 Laboratory procedure for calibration of digital multimeters | Multimeters/AC Voltmeters | Calibrator Wavetek 4808 |
| 4 | 100А-10А U> 65 ppm k =2 | LP-04-E-01 Laboratory procedure for calibration of digital multimeters | Multimeters/DC Ammeters | Calibrator Wavetek 4808/4600 |
| 5 | 10А-1000А U> 0.1 % k =2 | LP-04-E-01 Laboratory procedure for calibration of digital multimeters | Multimeters/DC Ammeters | Calibrator Wavetek 9100 |
| 6 | (100А-1А)@(100Hz-5kHz) U> 130 ppm k =2 | LP-04-E-01 Laboratory procedure for calibration of digital multimeters | Multimeters/AC Ammeters | Calibrator Wavetek 4808 |
| 7 | (1А-10А)@(100Hz-20kHz) U> 130 ppm k =2 | LP-04-E-01 Laboratory procedure for calibration of digital multimeters | Multimeters/AC Ammeters | Calibrator Wavetek 4808/4600 |
| 8 | (1A-20A)@(10Hz-10kHz) U> 1% k =2 | LP-04-E-01 Laboratory procedure for calibration of digital multimeters | Multimeters/AC Ammeters | Calibrator Wavetek 9100 |
| 9 | (20A-200A)@(10Hz-440Hz) U> 1.5% k =2 | LP-04-E-01 Laboratory procedure for calibration of digital multimeters | Multimeters/AC Ammeters | Calibrator Wavetek 9100 |
| 10 | (200A-1000A)@(10Hz-100Hz) U> 1% k =2 | LP-04-E-01 Laboratory procedure for calibration of digital multimeters | Multimeters/AC Ammeters | Calibrator Wavetek 9100 |
| 11 | 10Ω-100MΩ U> 20 ppm k =2 | LP-04-E-01 Laboratory procedure for calibration of digital multimeters | Multimeters/Ohmmeters | Calibrator Wavetek 4808 |
| 12 | 100MΩ – 400MΩ U> 0.8% k =2 | LP-04-E-01 Laboratory procedure for calibration of digital multimeters | Multimeters/Ohmmeters | Calibrator Wavetek 9100 |
| 13 | 100mV-1kV U> 10 ppm k =2 | LP-04-E-02 Laboratory procedure for calibration of calibrators | Multifunctional calibrator /DCV Calibrator | Digital multimeter Wavetek 1281 |
| 14 | (100mV-1V)@( 10Hz-100kHz) U> 140 ppm k =2 | LP-04-E-02 Laboratory procedure for calibration of calibrators | Multifunctional calibrator /ACV Calibrator | Digital multimeter Wavetek 1281 |
| 15 | (1V-100V)@( 10Hz-1MHz) U> 100 ppm k =2 | LP-04-E-02 Laboratory procedure for calibration of calibrators | Multifunctional calibrator /ACV Calibrator | Digital multimeter Wavetek 1281 |
| 16 | (100V-1kV)@( 10Hz-100kHz) – best uncertainty 200 ppm | LP-04-E-02 Laboratory procedure for calibration of calibrators | Multifunctional calibrator /ACV Calibrator | Digital multimeter Wavetek 1281 |
| 17 | 100А-1А U> 200 ppm k =2 | LP-04-E-02 Laboratory procedure for calibration of calibrators | Multifunctional calibrator /DCI Calibrator | Digital multimeter Wavetek 1281 |
| 18 | (100А-10А)@(40Hz-5kHz) U> 370 ppm k =2 | LP-04-E-02 Laboratory procedure for calibration of calibrators | Multifunctional calibrator /ACI Calibrator | Digital multimeter Wavetek 1281 |
| 19 | 10Ω-1GΩ U> 90 ppm k =2 | LP-04-E-02 Laboratory procedure for calibration of calibrators | Multifunctional calibrator /DC Resistance Calibrator | Digital multimeter Wavetek 1281 |